## Slide 1: Signal Acquisition and Data processing
NMR and MRI Application in Biotechnology

## Slide 2: LF NMR
LR-NMR technology is nondestructive and non-invasive, rapid, and does not require environmentally unfriendly solvents, as opposed to standard analytical methods.
Because of all of these, it is widely used for routine analysis in agriculture, food, polymer, petroleum, chemical, pharmaceutical and medical industries.

## Slide 3
Maran bench-top pulsed NMR analyzer (Resonance Instruments, Witney, United Kingdom) equipped with a permanent magnet and a 18 mm probe head operating at 23.4 MHz.
Signal Acquisition

## Slide 4
20 MHz minispec bench-top pulsed NMR analyzer (Bruker Analytic GmbH, Germany), equipped with a permanent magnet and a 10-mm temperature controlled probe head.
Signal Acquisition

## Slide 5
MiniEDU20 NMR/MRI system (Niumag, China) equipped with a permanent magnet and a 10 mm probe head operating at 0.08 T.
Signal Acquisition
Notes: Basic parameters:
1.Magnet type: permanent magnet
2.Magnetic field intensity: 0.5±0.08T
3.Probe coil: Ø10mm
4.Weight: 49.8Kg

## Slide 6
Free Induction Decay (FID)
A single RF pulse (normally a 90 degree pulse) is applied to the sample and the response is measured as a function of time after the RF excitation
Set up the NMR instrument.
Signal Acquisition

## Slide 7
CPMG Pulse Sequence
Named after the men who invented it (Carr, Purcell, Meiboom & Gill).
The CPMG sequence consists in applying a 90 degree RF pulse to the sample, followed by many 180 degree pulses.
Each time a 180 degree pulse is applied, the signal decay to the magnetic field is removed and a single data point is acquired.
The effects of the magnet on the signal are removed with this sequence
The decay constant of the exponential is called spin-spin or transverse relaxation time.
P90
P180
τ
P180
2τ
2τ
P180
P180
Signal Acquisition

## Slide 8
Inversion Recovery sequence
Spin-lattice or longitudinal relaxation
Consists of a 180 degree pulse (inversion pulse) and a 90 degree pulse.
Signal Acquisition

## Slide 9
Inversion Recovery- CPMG sequence
Contain T1 and T2 sequence.
Signal Acquisition

## Slide 10
System/Model
K
Signal f(t)
Noise ε(t)
Measurements s(t)
By proper discretization
Can be basically solved as a Least Squares problem
This LS problem is highly ill-conditioned. Therefore, it is not possible to be solved without a regularization method to improve its condition and the numerical stability of the solution procedure.
Signal Processing

## Slide 11
The most common regularization approaches apply the classical Tikhonov:
Tikhonov regularization approach gives a solution better posed and stable.
Nothing guarantees that the solution will be non-negative.
It tends to strongly over smooth the reconstructed spectrum function, which may result in merging close narrow peaks into a one wide peak and eliminate low volume peaks.
Another common practice for regularize the spectrum reconstruction problem is replacing the matrices K1 and K2 by their best low rank approximation using SDV compression.
Signal Processing

## Slide 12
A general formulation of the reconstruction problem, with l1 regularization and non-negativity constraint is indicated below:
F ≥ 0
In order to solve the equation, it was adjusted and used the optimization algorithm Primal-Dual Convex Objectives (PDCO)
Signal Processing

## Slide 13
(1) MR Signal collection –Relaxation Curve S
(2a) Data Processing –
ILT – 1D Mapping
(2b) Data Processing –
PDCO L1/L2 ITL –
2D T1-T2 Mapping
Increased peak generation and resolution
(3) Lipid Products Dictionary -
Time domain (TD) Peak Assignment
based on
Lipid Segmental Motion
4
3
2
1
(2a”) Improved ILT PDCO L1/L2 – 1D Mapping
* PDCO – Primal Dual Convex Optimization, Saunders, 2001
Key point: Choice of regulatory parameters L1 and L2
Sequence of Signal Analysis Developments with Ilt and PDCO

## Slide 14
Prof. Zeev Wiesman: wiesman@bgu.ac.il
Maysa Resende: mteixeiraresende@gmail.com
Dr. Janna Abramovich: janna@bgu.ac.il
Shosh Kravchick: kravch@bgu.ac.il
Plant Lipid Biotechnology Lab phone number: 074-7795272
Contact Information